## Buzz
Autoplate 5000 "Buzz" Preventive Maintenance Check Sheet
Quarterly Maintenance: Initial each activity when completed.
Month | Grease | Belts | Glycerol | Date | Validate | Date
2017-09-01
2017-12-01
2018-03-01
2018-06-01
2018-09-09
2018-12-18
2019-03-28
2019-06-01
2019-09-09
2019-12-18
2020-03-27
2020-06-01
2020-09-09
Annual Maintenance: Date and initial for each activity.
Date Scheduled | Date Performed | Activity |  | netID
2017-04-01 | 2017-06-08 | Replace stylus and tubing |  | sjr267
2017-06-01 | 2017-06-08 | Replace syringe |  | sjr267
2018-07-01 |  | Activities listed below
2019-07-01 |  | Activities listed below
2020-07-01 |  | Activities listed below
Annual Preventive Maintenance Activities:
-Replace the stylus tubing
-Replace the syringe
-Replace both quick connect fittings for both reagent bottles
-Replace both umbrella valves for both reagent bottles
-Replace the o-rings for all three pump supply valve hose barb adapters
-Replace all 1/16" ID silicone tubing:
*Tubing inside water bottle with ceramic weights - 12 cm
*Tubing inside disinfectant bottle with ceramic weights - 14.5 cm
*Pump supply valve to pump - 7.6 cm
*Water line - 75 cm
*Disinfectant line - 61 cm
Unscheduled Maintenance: date, netID, description, motivation
Date Performed | netID | Description |  | Motivation
2017-07-14 | sjr267 | Replaced the male component of both reagent bottle quick connects |  | Air was leaking into the system, and this was believed to be the source
2017-07-14 | sjr267 | Lubricated stylus boom ball bearings |  | Stylus was skipping over regions of the agar
2017-07-28 | sjr267 | Replaced the female component of both reagent bottle quick connects |  | Air was leaking into the system, and this was believed to be the source
2017-07-28 | sjr267 | Replaced the pump supply valve |  | The valve was open to both the water and the bleach when it was supposed to only be open to the water, resulting in bleach contamination of the water
2017-08-09 | sjr267 | Replaced all silicone tubing |  | Recommended by AI
2017-08-11 | sjr267 | Replaced o-ring for disinfectant hose barb on pump supply valve |  | It was missing
2017-08-11 | sjr267 | Replaced the isolation valve |  | Liquid was creeping up the stylus tip (~1 mm/10 sec) and second replicate deposition issue remained
Formulas:
Month = [Month, row above]+100

## Woody
Autoplate 4000 "Woody" Preventive Maintenance Check Sheet
Quarterly Maintenance: Initial each activity when completed.
Month | Grease | Belts | Glycerol | Date | Validate | Date
2017-09-01
2017-12-01
2018-03-01
2018-06-01
2018-09-09
2018-12-18
2019-03-28
2019-06-01
2019-09-09
2019-12-18
2020-03-27
2020-06-01
2020-09-09
Annual and Biennial Maintenance: Date and initial for each activity.
Date Performed | Activity |  | netID | Next Date to Perform
2016-03-09 | Replace stylus and syringe |  | sjr267 | March 2017 / March 2018
2017-08-11 | Replace stylus |  | sjr267 | August 2018
2017-08-11 | Replace pinch valve tubing |  | sjr267 | August 2018
Annual Preventive Maintenance Activities:
-Replace the stylus tubing
-Replace the pinch valve tubing:
*15 cm of 1/16" ID silicone tubing
Biennial Preventive Maintenance Activities:
-Replace the syringe
Unscheduled Maintenance: date, netID, description, motivation
Date Performed | netID | Description |  | Motivation
Formulas:
Month = [Month, row above]+100