Capnography

Capnometry= measurement and display of carbon dioxide concentration in the respiratory gases on a monitor

Max inspired and max expired CO2 concentrations are displayed

Capnography= graphic display of carbon dioxide concentration vs time or expired volume during a respiratory cycle

=carbon dioxide waveform or capnogram

Time capnograms are most common

Volume capnograms can be used to measure dead-space volume

Interpretation provides information about the patient and about the

equipment

Nondiverting and Diverting Monitors

Nondiverting Monitors= Mainstream Monitors- Measure carbon dioxide

concentration directly in the breathing system

Patient’s respiratory gas passes through a chamber with two windows

The sensor (light source and detector) fit over the chamber

The sensor also contains a heater to prevent water

condensation on the windows

Advantages: Fast response time, no requirement for scavenging gas,

ease of calibration (sealed chamber), use of few disposable items

Disadvantages: Necessity to place the sensor near the patient,

increase in apparatus dead space, potential for leaks, disconnection, obstruction, potential for the sensor to become dislodged from the chamber, exposure of the sensor to damage and measurement of carbon dioxide only

Diverting Monitors= Sidestream Monitors- Measure carbon dioxide

concentration in gas in the main unit

Sensor is located in the main unit, remote from the breathing system

Pump samples gas at a constant flow, via sampling tubing

Advantages: Minimal added dead space, lightweight patient interface,

potential for the measurement of multiple gases and possibility of use in places where the monitor needs to be remote from the patient

Disadvantages: Potential for sampling problems, delayed response, removal of gas at a given rate from the circuit, necessity to scavenge gas, potential for a change in gas composition, need for calibration, potential for gas mixing in the sampling line

Technologies

Infrared absorption

Most widely used technique

Only technique available in mainstream devices

Gases have 2+ dissimilar atoms in the molecule and therefore, have

unique and specific absorption spectra of infrared light (can be used for nitrous oxide and halogenated anesthetics)

Short warm-up time and quick response time--- inspired and expired concentrations

Water vapor must be removed from expired gas- absorbs infrared light

Mass spectrometry

Used uncommonly

Spreads gases and vapors of different molecular weights into a

spectrum, according to mass-to-charge ratios

Can be used to measure CO2, O2, nitrous oxide, and volatile anesthetics

Measures gases as concentrations as opposed to partial pressures

(assumes that sum of gases measured=100%)

Raman spectrometry

When laser light interacts with a gas molecule that has interatomic

bonds, some of its energy is converted into vibrational and rotational modes and a fraction of that energy is reemitted at various wavelengths characteristic of the molecule

Used in one monitor- Ra scal II- measures oxygen, nitrogen, carbon dioxide, nitrous oxide and up to 5 anesthetic agents

Fast response time, very fast startup

Requires little maintenance and is very accurate

Physiology

Normal capnogam can be divided into 5 phases

Phase IV=Early inspiration

Phase 0=Inspiration

No CO2 should be measured

Phase I= Early expiration

No CO2 should be measured (emptying of anatomical dead

space)

Phase II is a rapidly changing mixture of alveolar and dead-space gas

resulting in a steep increase in measured carbon dioxide

Phase III is the alveolar plateau, during which alveolar concentration

of carbon dioxide is measured resulting in a slightly increasing slope

Maximal concentration reached at the end of the plateau=

end tidal concentration

Alveolar carbon dioxide concentration in normal patients is only slightly lower than arterial carbon dioxide concentration (normal <5mmHg)

Arterial PCO2 and therefore end-tidal PCO2 result from the balance of carbon dioxide production and carbon dioxide elimination (alveolar ventilation)

Because CO2 is the major factor controlling breathing any change in CO2

production will induce a proportional change in alveolar ventilation so that PCO2 remains constant

Exception- patients under GA or those with respiratory, muscular,

neurologic disease

Increased dead space ventilation (i.e. ventilation of alveoli that are not

perfused) increases end-tidal CO2 and arterial PCO2 gradient

V/Q scattering and decreased pulmonary blood flow (PTE, low cardiac

output)

End-tidal CO2 cannot be used to estimate arterial PCO2

Capnogram Interpretation

Non-invasive, continuous monitoring

Equipment

Increased apparatus dead space, rebreathing with a circle system

(faulty unidirectional valve, exhausted CO2 absorbent), or rebreathing with a nonbreathing system (inadequate fresh gas flow, leak or disconnection of the inner tubing of the Bain circuit)

Elevated baseline i.e. increased inspired CO2 (and if vent doesn’t change increase in expired CO2)

End-tidal to arterial PCO2 gradient will be normal or decreased

Slope of phase IV usually decreased

Obstruction to expiration

Decrease in the slope of phase II, increase in the slope of phase

III, decrease in end-tidal PCO2 (if inspiration starts before expiration complete)

Inadequate endotracheal cuff seal

Low, non-0 end-tidal readings

Patient

Changes in patient carbon dioxide production without associated

changes in alveolar ventilation will result in changes in end-tidal CO2 with a normal waveform, no inspired CO2 and a normal end tidal to arterial PCO2 gradient

Apnea results in absence of a waveform

Hyperventilation results in decreased end-tidal PCO2, otherwise normal waveform, no inspired CO2 and a normal ET to arterial gradient

Hypoventilation results in opposite changes

Upper airway obstruction results in changes similar to equipment obstruction

Decreased transport of CO2 to the lungs (i.e. decreased CO, PTE) will result in decreased end-tidal PCO2, normal capnogram, no inspired CO2, increased end-tidal to arterial PCO2 gradient

Cardiac arrest results in a sudden decrease of end-tidal PCO2 to 0 or near 0 values; can be used as indicator of success of CPR

Bronchoconstriction can result in prolonged expiration (phases II and III), decreased slope of phase II, increased slope of phase III

Cardiogenic oscillation sometimes seen on capnogram- small oscillations in phase IV; slope of downstroke also decreased

Frequency of oscillations corresponds to HR

Curare cleft= decrease in CO2 during the alveolar plateau

Due to spontaneous inspiratory effort in mechanically ventilated patients

Current Literature:
Agreement between values for arterial and end-tidal partial pressures of carbon dioxide in spontaneously breathing, critically ill dogs

Efrat Kelmer, dvm, ms, dacvecc; Lindsey C. Scanson; Ann Reed, ms; Lydia C. Love, dvm

JAVMA 2009

Objective—To determine agreement between arterial partial pressures of carbon dioxide (Paco2) and end-tidal partial pressures of carbon dioxide (Petco2) measured with a nasal catheter in spontaneously breathing, critically ill dogs.

Design—Validation study.

Animals—26 client-owned dogs admitted to an intensive care unit for various conditions.

Procedures—Paco2 was measured with a commercial blood gas analyzer, and Petco2 was measured with a sidestream capnograph attached to a nasal catheter. Measurements were obtained twice (ie, with and without supplemental oxygen). Paired values were compared by means of the Pearson correlation method. Level of agreement was assessed by means of the Bland-Altman method.

Results—Mean difference between Paco2 and Petco2 when dogs did not receive supplemen- tal oxygen (mean ± SD, 3.95 ± 4.92 mm Hg) was significantly lower than mean difference when dogs did receive supplemental oxygen (6.87 ± 6.42 mm Hg). Mean difference in dogs with a condition affecting the respiratory system (8.55 ± 5.43 mm Hg) was significantly higher than mean difference in dogs without respiratory tract disease (3.28 ± 3.23 mm Hg). There was a significant linear correlation and good agreement between measured values of Paco2 and Petco2. Catheter size, ventilatory status, and outcome were not significantly associated with mean difference between Paco2 and Petco2.

Conclusions and Clinical Relevance—Results suggested that nasal capnography is a clinically relevant method of estimating Paco2 in spontaneously breathing, critically ill dogs, but that values should be interpreted with caution in dogs receiving supplemental oxygen and in dogs with conditions affecting the respiratory system.