Common components of all anesthetic units:

- source of oxygen

- regulator for oxygen (this may be part of the gas supply system)

- flowmeter for oxygen

- vaporizer

Used with a breathing circuit and anesthetic waste gas scavenging system for delivery of anesthetic to the patient

Anesthetic machines have 2 gas supplies:

- small, high‐pressure tanks attached directly to the machine

- hospital’s central pipeline supply

Modern anesthetic machine

Gas flow

- Anesthesia delivery apparatus: (1) gas delivery system, (2) vaporizer, (3) breathing circuit, (4) waste gas scavenge system

- High‐, intermediate‐, and low‐pressure areas in the anesthetic machine

Flowmeter

- Controls the rate of gas delivery to the low‐pressure area of the anesthetic machine and determines the fresh gas flow (FGF) to the anesthetic circuit

- Flowmeters are gas specific and calibrated at 760 mmHg and 20C

Vaporizers

- Vaporizers change liquid anesthetic into vapor and meter the amount of vapor leaving the vaporizer

- They work by splitting the carrier gas to flow into the vaporizing chamber (where it picks up anesthetic vapor) or to the bypass channel where it does not

- Factors that alter vaporizer output: temperature, flow, and pressure

- Most modern precision compensated vaporizers will maintain consistent output at flows between 0.5 and 10 L/min and temperatures between 15 and 35C

- Three main styles of vaporizers: (1) Ohmeda Tec, (2) Drager Vapor, (3) Penlon Sigma series

Use of the wrong anesthetic in an agent‐specific vaporizer

- Results in lower or higher anesthetic output (as a result of varying vapor pressures)

Oxygen flush valve

- Flush valves are designed to rapidly deliver large volumes of non‐anesthetic containing gas to the patient breathing circuit in emergency situations

- Bypasses the flowmeter and vaporizer

Common gas outlet

- Gas reaching the common gas outlet has traveled from the gas supply (cylinder or pipeline), through the regulator, flowmeter, and vaporizer

Breathing system

- Primary purposes: (1) direct oxygen to the patient, (2) deliver anesthetic gas to the patient, (3) remove carbon dioxide from inhaled breaths (or prevent significant rebreathing of carbon dioxide), (4) provide a means of controlling ventilation.

- Two groups: those designed for rebreathing of exhaled gases (rebreathing or partial rebreathing system) and those designed to be used under circumstances of minimal to no rebreathing (non‐rebreathing systems)

Rebreathing (Circle systems)

    - Designed to produce a unidirectional flow of gas through the system and has a means of absorbing CO2 from exhaled gases

    - Components include: fresh gas inlet, inspiratory one‐way valve, breathing tubes, expiratory one‐way valve, APL valve, reservoir bag, and carbon dioxide absorber

      - Full rebreathing: circle system using flow rates equal to, or nearing, the metabolic oxygen consumption of the patient, between 3 and 14 mL/kg/min

      - Partial rebreathing:  circle system using a flow rate greater than metabolic oxygen consumption (e.g., 20 mL/kg/min) but less than that required to prevent rebreathing.

      - Non- (minimal) rebreathing: circle system using flow rates greater than 200 mL/ kg/min (flow rates that would normally not be used in most circumstances). For use in very small patients that need high floe rates (<5kg).

Questions

- Name three factors that can affect vaporizer pressure.

- True or False. Partial rebreathing is a circle system that uses flow rates equal to, or nearing, the metabolic oxygen consumption.

- A green H medical gas cylinder contains how many liters of gas?

  - 660L of oxygen

  - 6900L of oxygen

  - 1590L of Nitrous oxide

  - 15800L of nitrous oxide

- Name the 4 common components to all anesthetic units

Answers

- Name three factors that can affect vaporizer pressure

Temperature, flow, and pressure

- True or False. Partial rebreathing is a circle system that uses flow rates equal to, or nearing, the metabolic oxygen consumption

      - Full rebreathing: circle system using flow rates equal to, or nearing, the metabolic oxygen consumption of the patient, between 3 and 14 mL/kg/min

      - Partial rebreathing:  circle system using a flow rate greater than metabolic oxygen consumption (e.g., 20 mL/kg/min) but less than that required to prevent rebreathing.

      - Non- (minimal) rebreathing: circle system using flow rates greater than 200 mL/ kg/min (flow rates that would normally not be used in most circumstances). For use in very small patients that need high floe rates (<5kg).

- A green H medical gas cylinder contains how many liters of gas?

  - 660L of oxygen

  - 6900L of oxygen

  - 1590L of Nitrous oxide

  - 15800L of nitrous oxide

- Name the 4 common components to all anesthetic units

- source of oxygen

- regulator for oxygen (this may be part of the gas supply system)

- flowmeter for oxygen

- vaporizer

Used in conjunction with a breathing circuit and anesthetic waste gas scavenging system for delivery of anesthetic to the patient