Key Points

- Oxygen therapy

- ETT:  high-volume, low-pressure cuffs are preferred to minimize the risk of ischemic tracheal injury – the pressure exerted by the cuff on the tracheal wall is similar to the intracuff pressure.Tracheal wall pressures > 48 mmg may impede capillary blood flow ( ischmiec tracheal damage); < 18 mmHg may potentially increase the risk of aspiration

  - Cuff monitor to inflate high-volume, low-pressure cuffs

  - Leak test (inflate the cuff until a leak is no longer audible while maintaining AW pressure of 20-30 cmH2O)

- Supraglottic airway devices

  - Simpler, faster to place than ETT in some species

  - Do not require laryngoscope

  - Do not enter the larynx or trachea

  - Form a seal around the glottis

  - Does not produce greater leakage of anesthetic gases

Lumb and Jones – Chapter 3 Anesthesia Equipment

Endotracheal tubes

- Cuffed Murphy-type

- Magill-type (no Murphy hole)

- Cole-type

- Guarded tube

- Self-sealing endotracheal tube

- High-volume, low pressure cuffs are preferred

  - Pressure exerted on tracheal wall by the cuff approximately = intracuff pressure

  - Tracheal wall pressure > 48 mmHg  decreased capillary blood flow; < 18 mmHg  risk of aspiration

  - Cuff monitor to inflate cuff or leak test (inflating the cuff until a leak is no longer audible while maintaining AW pressures of 20-30 cmH2O)

- Endotracheal tubes for isolating one lung

  - Indications:  improving surgical conditions for various thoracic procedures (thoracoscopy), control of contamination or hemorrhage, when unilateral pathology…

    - Double lumen tube

    - Bronchial blocker

    - Long standard ETT as an endobronchial tube

Supraglottic airway devices

- Aka laryngeal mask airways

- May be simpler and faster to place than ETT in some species

- Do not require use of a laryngoscope

- Do not enter the larynx or trachea

- Form a seal around the glottis

- Not associated with greater leakage of anesthetic gases

Laryngoscopes

- MacIntosh – curved blade

- Miller – straight blade

Oxygen administration

- If patient objects to mask so vigorously, the increase in oxygen consumption associated with restrain may nullify the benefits of greater FiO2

- With a tight fitting mask, higher flow rates of oxygen tend to produce greater FiO2 values and less rebreathing of expired carbon dioxide

- A loose-fitting mask that allows room air to be entrained can be used but will reduce FiO2

- Nasal insufflation – relatively high flow rate; patient inspires oxygen and room air: the relative proportions of each being determined primarily by the oxygen flow rate and the rate of gas flow during inspiration

  - Tip positioned in the nasopharynx

  - Tube should ideally have several fenestrations to minimize jetting lesions from developing in the nasopharyngeal mucosa

  - Changing the catheter to the opposite nasal passage every 1-2 days have beend recommended to prevent pressure necrosis, jet lesions, and accumulation of mucus

  - Bilateral nasal insufflation – up to 80% FiO2 at 200 ml/kg/min

- Tracheal insufflation

  - Catheter through the nares into the trachea

- Oxygen cages

  - 30-60% O2; may require flow as high as 15 L/min

  - For small animals: recommendations –

    - Flow rates of oxygen: < 10 L/min

    - Temperature: 22°C (71.6°F)

    - Cage humidity: 40-50%

- Oxygen toxicity

  - Endothelial damages, destruction of alveolar cells, increased microvascular permeability  edema, hemorrhage, congestion; alveolar type II and fibroblast proliferation  fibrosis

    - Length of time that a patient’s PaO2 is elevated may be MORE predictive of oxygen toxicity than the duration of exposure to high FiO2

    - Guidelines: PaO2 of 70 mmHg as an endpoint of oxygen therapy, use lowest FiO2 possible to achieve PaO2 of 70 and do not use FiO2 >0.6 for longer than 24h

Questions

- Name oxygen supplementation techniques the maximal FiO2 that can be reached with each technique

- Describe how to do a “leak test”.

- Which of the following patient is the most likely to suffer from the side effects of oxygen toxicity

  - Patient A: FiO2 = 40% and PaO2 = 130 mmHg for 10 hours

  - Patient B: FiO2 = 65% and PaO2 = 77 mmHg for 10 hours

- Name alternative techniques that could be use in the case of difficult ETT placement

Questions & Answers

- Name oxygen supplementation techniques and the maximal FiO2 that can be reached with each technique

Oxygen chamber: 50-60%

Nasal insufflation: 70% (up to 80% with bilateral nasal insufflation)

Flow-by oxygen: 40%

Face mask: 60%

Tracheal insufflation: 60%

- Describe how to do a “leak test”.

Place ETT and connect ETT to breathing tube

Connect a syringe filled with air to the pilot balloon and inflate the cuff while listening for a leak

At the same time an assistant applies and maintains AW pressures of 20-30 cmH2O by compressing the reservoir bag (pop-off valve is closed)

Stop inflating the cuff when a leak is no longer audible

- Which of the following patient is the most likely to suffer from the side effects of oxygen toxicity

  - Patient A: FiO2 = 40% and PaO2 = 130 mmHg for 10 hours

  - Patient B: FiO2 = 65% and PaO2 = 77 mmHg for 10 hours

Length of time that a patient’s PaO2 is elevated may be MORE predictive of oxygen toxicity than the duration of exposure to high FiO2

- Name alternative techniques that could be use in the case of difficult/impossible ETT placement

Supraglottic airway devices

Tracheostomy

Nasotracheal intubation

Wire- or tube-guided techniques (antegrade or retrograde)

Endoscope-guided techniques