Take-home messages

- Two main types of ventilators are compressed gas powered (two-circuit systems) that have bellows and electronic piston (one-circuit systems)

- Main mechanisms of alterations between desired tidal volume and actual delivered tidal volume include patient and system compliance, leaks, and fresh gas flow (mostly for smaller patients)

- Bag size for a patient not on mechanical ventilation should be approximately 50-200mls/kg

Anesthetic equipment

- Reservoir bag

  - Provides compliant area for gas for patient

  - Should be 50-200ml/kg

    - If bag is too large, there may be a delay in changes in anesthetic gas concentration and patient plasma concentration or difficult to see breathing movements in bag

- Adjustable pressure-limiting valve (pop-off valve)

  - Allows gas to escape circuit

- CO2 absorber

  - Consist of a base that allows absorption and salt formation with acid (CO2)

Ventilators

- Major Control Variable

  - Pressure or volume control

- Driving force

  - Electric powered piston

    - “Single circuit ventilator”

  - Compressed gas powered

    - “Dual circuit ventilator”

    - Two chambers exist, one that fills with compressed gas (usually 100% O2) surrounding a bellows that gets compressed and delivers gas to patient

- Factors that affect delivered tidal volume (compared to set tidal volume)

  - Fresh gas flow

    - If set too high in very small patients, can contribute to tidal volume delivered to patient

  - Compliance and compression volumes

    - Increased system compliance will decrease amount of gas delivered to patient

    - Increased patient compliance will increase amount of gas delivered to patient

  - Leaks in system

    - Around ET tube

    - Side-stream gas flow monitors (NICO)

    - Pneumothorax

    - Etc…

Questions

- Name 3 factors that can cause a difference between preset tidal volume and actual tidal volume when using a mechanical ventilator.  Describe the difference (i.e. Patient’s tidal volume will be greater/less than expected tidal volume)

- Name 2 downsides to having a reservoir bag that is >200mls/kg

Answers

- Name 3 factors that can cause a difference between preset tidal volume and actual tidal volume when using a mechanical ventilator.  Describe the difference (i.e. Patient’s tidal volume will be greater/less than expected tidal volume)

  - Leak (will cause actual tidal volume to be lower than set volume)

  - Increased machine compliance (will cause actual tidal volume to lower than set)

  - Increased patient compliance (will cause actual tidal volume to be higher than set)

  - Increased fresh gas flow rate (will cause actual tidal volume to be higher than set)

- Name 2 downsides to having a reservoir bag that is >200mls/kg

  - Cannot see movement in bag when patient breaths

  - Delay in changes of vaporizer settings to % anesthetic delivered